Manoj Agarwal,

Sharad Pandey,

Pankaj Kumar,

L.N. Gupta

Shunt in scrotum: unusual shunt complication in an operated case of post TBM hydrocephalus

DOI: 10.33962/roneuro-2021-037

Shunt in scrotum: unusual shunt complication in an operated case of post TBM hydrocephalus

Manoj Agarwal, Sharad Pandey, Pankaj Kumar, L.N. Gupta1

Department of Neurosurgery, A.B.V.I.M.S. and Dr Ram Manohar Lohia Hospital, New Delhi, India

Keywords

Ventriculo Peritoneal Shunt (VPS),

complications,

distal catheter migration, scrotum

Corresponding author:

Sharad Pandey

Department of Neurosurgery, A.B.V.I.M.S. and Dr Ram Manohar Lohia Hospital, New Delhi, India

drsharad23@yahoo.com

Copyright and usage. This is an Open Access article, distributed under the terms of the Creative Commons Attribution Non–Commercial No Derivatives License (https://creativecommons .org/licenses/by-nc-nd/4.0/) which permits non-commercial re-use, distribution, and reproduction in any medium, provided the original work is unaltered and is properly cited.

The written permission of the Romanian Society of Neurosurgery must be obtained for commercial re-use or in order to create a derivative work.

ISSN online 2344-4959

© Romanian Society of Neurosurgery

First published

June 2021 by

London Academic Publishing

www.lapub.co.uk

Abstract

The role of shunt placement is to divert cerebrospinal fluid (CSF) from within the ventricles to an alternative location most commonly peritoneal space. Ventriculo Peritoneal Shunt (VPS) is associated with many complications viz over drainage, valve failure, breaking of catheter, catheter obstruction, coiling of catheter, spontaneous knot formation, infection and migration of distal catheter and all of them finally leading to obstructive hydrocephalus. One such complication is distal catheter migration to a rare but possible site i.e. scrotum. We reported a case of scrotal migration of shunt as a late complication of VPS insertion.

Background

The role of shunt placement is to divert cerebrospinal fluid (CSF) from within the ventricles to an alternative location (most commonly peritoneal space, rarely to atria, pleural space) and from lumbar space to peritoneal cavity.

Ventriculo Peritoneal Shunt (VPS) is associated with many complications namely over drainage, valve failure, breaking of catheter, catheter obstruction, coiling of catheter, spontaneous knot formation, infection and migration of distal catheter and all of them finally leading to obstructive hydrocephalus [1,2]. One of the rare sites of distal catheter migration is scrotum. Scrotal migration of distal shunt catheter is more common in children due to patent processus vaginalis [3]. We reported a case of scrotal migration of shunt as a late complication of VPS insertion.

Case presentation

An eighteen months old male child, a follow up case of tubercular meningitis (TBM) with hydrocephalus was operated by right VPMP shunt via right Keens point in January 2020 at our tertiary care centre. Post operatively patient's GCS improved and patient was discharged. After approximately 03 months patient presented with gradual increase in size of bilateral scrotum in the past 2 months, intermittent

fever and on & off vomiting and was admitted (Fig.1). On local examination cough impulse was present in the right inguinal region & a tube like structure was palpable in the right hemiscrotum, external genitalia was normal. Neurological examination was normal with no focal neurological deficit, GCS was E4V5M6, head circumference was normal for age, no signs of shunt obstruction were present.

Investigations

X-ray of neck, chest, abdomen & pelvis were conducted which showed an intact VP shunt with extension of the peritoneal end into the right hemi scrotum. USG of inguinoscrotal region was also suggestive of same.

Treatment

Surgical intervention was planned in this rare case. Re-exploration of the abdominal wound was done, the peritoneal end of the shunt was identified and it was pulled and repositioned in peritoneal cavity under C-ARM guidance and the abdominal wound was closed in layers. Post-operative X-ray of chest, abdomen & pelvis showed normal position of peritoneal end of VP shunt (Fig.2).

Outcome and follow up

The child was kept under observation and discharged after 48 hours since the postoperative period was uneventful and was advised for follow up in general Surgery OPD and Neurosurgery OPD.

Discussion

The most common surgery performed worldwide for hydrocephalus is VP shunt [3]. Shunting is performed to divert the CSF from dilated ventricles to peritoneal cavity from where CSF is absorbed. Apart from the morbidity associated with hydrocephalus, shunt surgeries have their own complications which are well known [2-6].The reported incidence of shunt-related abdominal complications is 10-30%[7] Various authors have reported migration of the distal end of peritoneal catheter into the scrotum through a patent processus vaginalis[8-11]. The processus vaginalis may remain patent until 1 year of life in 50-60% cases and until 2 years of life in 40-50% of cases [3]. Rowe et al[12] have explained the aetiology of peritoneal catheter migration, which has been universally accepted over the years by various authors[13]. Peritoneal cavity distension due to draining CSF also prevents the obliteration of processus vaginalis which had also been reported by Ho et al [14] and Ozveren et al[15]. In the present case, a male child had a patent processus vaginalis which was either congenital or because of peritoneal CSF drainage. In our case, right-sided VP shunt was performed and it would be difficult to comment and is another research study topic whether the side has something to do with the shunt migration as there is no such report in the literature comparing the right and left patent processus vaginalis. We used Chabbra medium pressure VP shunt in our case. Shunt migration can occur either as an early or a late complication. Ozveren et al[15] have reported shunt migration within 24 h, while the average length of time reported in other series was 6.8 months. In our case, the time interval was approximately 3 months. Thus this kind of shunt complication needs to be reported and should be kept into mind in dealing with cases who had undergone shunt insertion surgery.

Scrotal VP shunt migration has been reported in 26 children, and the time of presentation varies from 1 month to 2.5 years (mean 3.8 months, median 1 month) after shunting, with 21 out of 26 migrations occurring within 6 months, 88% involving the right side [16].

Ramareddy RS describe the spontaneous resolutionof hydrocele and retreat of migrated peritoneal portion of VP shunt, and also discuss the pathogenesis and management [16].

References

- Borkar SA, Satyarthee GD, Khan RN, et al. Spontaneous extrusion of migrated ventriculoperitoneal shunt catheter through chest wall: a case report. Turk Neurosurg 2008;18:95–8.

- Agha PP, Amendola MA, Shirazi KK, et al. Abdominal complications of ventriculoperitoneal shunts with emphasis on the role of imaging methods. Surg Gynecol Obstet 1983;156:473–7.

- Grosfeld IL, Cooney DR. Inguinal hernia after ventriculoperitoneal shunt for hydrocephalus. J Pediatr Surg 1974;9:311–15.

- Patel CD, Matloub H. Vaginal perforation as a complication of ventriculoperitoneal shunt. J Neurosurg 1973;38:761–2.

- Redman JF, Seibert JJ. Abdominal and genitourinary complications following ventriculoperitoneal shunts. J Urol 1977;119:295–7.

- Wakai S. Extrusion of a peritoneal catheter through the abdominal wall in an infant. Case report. J Neurosurg 1982;57:148–9.

- Bryant MS, Bremer AM. Abdominal complications of ventriculoperitoneal shunts: case reports and review of literature. Am Surg 1988;54:50–5.

- Bristow DL, Buntain WL, James HL. Ventriculoperitoneal (VP) shunt migration causing an acute scrotum: a case report of Doppler evaluation. J Pediatr Surg 1978;13:538–9.

- Fuwa I, Matsukado Y, Itoyama Y, et al. Migration of a dissected peritoneal shunt catheter into the scrotum. Brain Dev 1984;6:336–8.

- Levy SH, Cooper P, Shiffman D. Simulated testicular torsion in a neonate: complication of ventriculoperitoneal shunt. Urology 1977;9:174–6.

- Ramani PS. Extrusion of abdominal catheter of ventriculoperitoneal shunt into the scrotum. J Neurosurg 1974;40:772–3.

- Rowe MI, Copelson LW, Clatworthy HW. The patent processus vaginalis and inguinal hernia. J Pediatr Surg 1969;4:102–7.

- Kita D, Hayashi Y, Kinoshita M, et al. Scrotal migration of the peritoneal catheter of a ventriculoperitoneal shunt in a 5-year-old male: case report. Neurol Med Chir (Tokyo) 2010;50:1122–5.

- Ho CC, Jamaludin WJ, Goh EH, et al. Scrotal mass: a rare complication of ventriculoperitoneal shunt. Acta Medica (Hradec Kralove) 2011;54:81–2.

- Ozveren MF, Kazez A, Cetin H, et al. Migration of the abdominal catheter of a ventriculoperitoneal shunt into the scrotum—case report. Neurol Med Chir (Tokyo) 1999;39:313–15.

- Ramareddy RS , Alladi A . Scrotal migration of ventriculoperitoneal catheter and hydrocele resolving spontaneously . Indian J Neurosurg 2012;1:172-4.